## Introduction
Technical specifications for selection of essential in vitro diagnostics (WHO EDL)
Introduction
Technical specifications enable procurement officers and agencies to adopt a systematic, unbiased approach to the selection of medical products, resulting in procurement, leasing or donation of appropriate commodities in line with the requirements of each level of the health-care system and local setting. Technical specifications for in vitro diagnostics (IVDs) constitute a set of criteria and baseline requirements to ensure good quality, safety, performance and efficacy. They are a decisive companion to the WHO Model list of Essential In Vitro Diagnostics (WHO-EDL) that will allow Member States, donor agencies and nongovernmental organizations to select products within each test category of the WHO EDL and will guide procurement decisions. The present publication defines the basic generic technical characteristics of IVDs listed in the WHO EDL. WHO will review and update this document regularly.
Objective
To provide information to support decisions on procurement of good-quality, safe in vitro diagnostics to improve in vitro diagnostic testing.
Target audience
This tool is intended for programme managers, laboratory managers, procurement officers and planning officers in ministries of health. It will also inform biomedical engineering professionals, private health sector, medical device industry and intergovernmental and international agencies about the characteristics of the products required.
Methods
The method used to prepare the technical specifications consisted of a review of existing WHO guidance and specifications for the selection, procurement, donation or lease of good-quality IVDs and medical devices in order to prepare a common template of generic specifications. The guidance documents analysed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (1), WHO Guidelines on hepatitis B and C testing (2), the Third WHO Model List of Essential In Vitro Diagnostics (3) and WHO internal procurement templates.
In the second phase of the project, WHO commissioned an external team of consultants to develop technical specifications for the IVDs listed in the 3rd WHO-EDL. The process included: a market survey and a horizon scan of existing tests and platforms; analysis of products in the market after approval by regulatory agencies; analysis of international, regional and country standards; review of WHO guidance and guidelines, review of peer-reviewed publications, textbooks and published qualitative and quantitative comparisons of tests and platforms; comparison of test specifications (i.e. accuracy, precision, positive and negative predictive value, likelihood ratios) from catalogues and previous tenders by United Nations agencies with generic formulations; and data from external quality assessments and technical expert opinion. Each template was drafted and reviewed by at least two qualified technical experts. The data presented represents a universal template, generally applicable to a broad range of analysers and methods. All the consultants provided declarations of conflicts of Interest, and no conflicts were found. The draft technical specifications were reviewed internally by experts before being posted on the WHO website and shared with relevant networks for public consultation. The comments received were reviewed and included as appropriate, and the technical specifications were finalized.
References
1. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment, Geneva: World Health Organization; 2017 (https://apps.who.int/iris/handle/10665/255577). 2. Guidelines on hepatitis B and C testing. Geneva: World Health Organization; 2017 (https://www.who.int/publications/i/item/9789241549981). 3. Third WHO Model List of Essential In Vitro Diagnostics, Geneva:World Health Organization; 2021 (https://www.who.int/publications/i/item/9789240019102).

## Content
Technical specifications for selection of essential in vitro diagnostics (WHO EDL)
The present tool includes technical specifications for 16 in vitro diagnostic tests for HBV listed in th WHO EDL
Each technical specification is presented in a standard template, which is organized as follows: name of the IVD test or laboratory equipment at the top, and 88 characteristics of technical specifications organized into 13 categories: Name, category and coding; Intended use; Performance characteristics; Technical and operational characteristics; Instrument physical and technical characteristics; Infrastructure requirements; Accessories, consumables, spare parts and other components; Documentation; Environmental and safety requirements; Training, installation and use; Warranty and maintenance; Decommissioning; Quality and registration.
In order to review technical specifications for a specific IVD category, go to the test menu tab, and click on the name of the relevant diagnostic test category. The selected item will be displayed on the screen. Scroll down to review the basic generic technical characteristics of the item.

## Tests menu
Technical specifications for selection of essential in vitro diagnostics: WHO EDL
No. | Page in TRS 1031/EDL 3 | Diagnostic test name | Test purpose | Assay format
1 | P281 for no laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen for HBV infection or to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults | Rapid diagnostic test (RDT)
2 | P325 for clinical laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen for HBV infection and to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults | Rapid diagnostic test (RDT)
3 | P325 for clinical laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen for HBV infection and to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults | Immunoassay Manual method
4 (Auto) | P325 for clinical laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen for HBV infection and to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults | Immunoassay Automated method
5 | P325 for clinical laboratory | Quantitative HBV nucleic acid test | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment | Nucleic acid test (NAT) Manual method
6 (Semi auto) | P325 for clinical laboratory | Quantitative HBV nucleic acid test | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment | Nucleic acid test (NAT) Semi automated method
7 (Fully auto) | P325 for clinical laboratory | Quantitative HBV nucleic acid test | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment | Nucleic acid test (NAT) Fully automated method
8 | P325 for clinical laboratory | Hepatitis B e antigen (HBeAg) | To stage chronic HBV infection and to determine the need for treatment | Immunoassay Manual method
9 (Auto) | P325 for clinical laboratory | Hepatitis B e antigen (HBeAg) | To stage chronic HBV infection and to determine the need for treatment | Immunoassay Automated method
10 | P326 for clinical laboratory | IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) | To aid in the diagnosis of acute HBV infection in the context of an outbreak investigation | Immunoassay Manual method
11 (Auto) | P326 for clinical laboratory | IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) | To aid in the diagnosis of acute HBV infection in the context of an outbreak investigation | Immunoassay Automated method
12 | P326 for clinical laboratory | Antibodies to hepatitis B surface antigen (anti-HBs) | To determine immune status after HBV vaccination | Immunoassay Manual method
13 (Auto) | P326 for clinical laboratory | Antibodies to hepatitis B surface antigen (anti-HBs) | To determine immune status after HBV vaccination | Immunoassay Automated method
14 | P346 for blood screening laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen blood donations for HBV | Rapid diagnostic test (RDT)
15 | P346 for blood screening laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen blood donations for HBV | Immunoassay Manual method
16 (Auto) | P346 for blood screening laboratory | Hepatitis B virus surface antigen (HBsAg) | To screen blood donations for HBV | Immunoassay Automated method
Excel cells coloured in ivory are for disease-specific IVDs for use in community settings and health facilities without laboratories.
Excel cells coloured in green are for disease-specific IVDs for use in clinical laboratories.
Excel cells coloured in pink are for disease-specific IVDs for blood-screening laboratories.

## 1
Hepatitis B surface antigen (HBsAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus surface antigen (HBsAg)
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for hepatitis B virus surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B virus surface antigen, HBsAg, hepatitis, screening, rapid diagnostic test, RDT
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen for HBV infection or to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Capillary whole blood, venous whole blood
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained health-care providers, trained lay and peer workers
13 | Level of the health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 97%
15 | Clinical specificity | ≥ 98%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | < 2 IU/mL
18 | Rate of invalid, error and unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Lateral flow assay
22 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. Specimens can be stored at 2–8 °C. Do not freeze whole-blood specimens.
23 | Specimen(s) volume | 25–100 µL
24 | Type of result | Qualitative (reactive or non-reactive)
25 | Time to result | 15–30 min
26 | End-point stability | 5–10 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Maximum, 2 steps
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 2 specimens per hour per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | A built-in quality control line that must appear when a specimen is added is part of internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the procedure in case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight. Any specific requirements once the reagents are opened should be clearly stated.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents after manufacture (months) | ≥ 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 of Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/Test menu availability | Not applicable
Infrastructure requirements
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
63 | Kit component – test/reagents/consumables (if relevant) | Test cards or test cassettes or test strips
64 | Reagent/kit size (number of tests) | ≥ 20 tests
65 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipettes are not provided, tips for micropipettes are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettes for venous whole blood are not provided, tips for micropipettes are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick or capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed, more than one timer would be an advantage. If disposable pipettors are not provided, a calibrated micropipette and a micropipette stand are required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. Instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre, with the dates on which the lot was used.
71 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard related to handling, storage or use of the product; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This is usually described in the material hazards data sheet
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacturer was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency, and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay in a RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 2
Hepatitis B virus surface antigen (HBsAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus surface antigen (HBsAg)
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for hepatitis B virus surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B virus surface antigen, HBsAg, hepatitis, screening, rapid diagnostic test
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen for HBV infection, or to aid in the diagnosis of chronic and acute HBV infection: children, adolescents and adults
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 97%
15 | Clinical specificity | ≥ 98%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant impact on the performance of the product.
17 | Limit of detection | < 2 IU/mL
18 | Rate of invalid, error and unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Lateral flow assay
22 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. Specimen can be stored at 2–8 °C. Do not freeze whole-blood specimens.
23 | Specimen(s) volume | 25–100 µL
24 | Type of result | Qualitative (reactive or non-reactive)
25 | Time to result | 15–30 min
26 | End-point stability | 5–10 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Maximum of two steps
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 2 specimens per hour per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | A built-in quality control line that must appear when a specimen is added plays a role of Internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the case of invalid results. The package insert should state how positive, negative or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight. Any specific requirement once the reagents are opened should be clearly stated.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature or usage of centrifuge will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | Test cards or test cassettes or test strips
64 | Reagent/kit size (number of tests) | ≥ 20 tests
65 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipettes are not provided, tips for micropipettes are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettes for venous whole blood are not provided, tips for micropipettes are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified for the kits/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually included in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | An available local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay in RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions.Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 3
Hepatitis B virus surface antigen (HBsAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus surface antigen (HBsAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay for hepatitis B virus surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B surface antigen, HBsAg, hepatitis, microtitre plate enzyme immunoassay, HBV
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen for acute and chronic hepatitis B virus (HBV) infection: children, adolescents and adults
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 99%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.2 IU/mL
18 | Rate of invalid, error and unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 11%; inter-assay precision (reproducibility): CV ≤ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Enzyme immunoassay, microwells coated with antibodies to HBsAg
22 | Specimen(s) stability | At 2–8 °C for 72 h and at –20 °C beyond 72 h
23 | Specimen(s) volume | ≤ 100 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 240 min
26 | End-point stability | 10–20 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. The proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Specific requirements once the reagents are opened should be clearly stated.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C beyond 72 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C beyond 72 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C beyond 72 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirements once the reagents are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassys and enzyme-linked immunosorbent assays. See the instruction for use or package insert of the kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | Coated plate with wells, conjugate and HBsAg positive and negative control sera and package insert/instructions for use.
64 | Reagent/kit size (number of tests) | ≧ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader, enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually contained in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. There is a WHO prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 4(Auto)
Hepatitis B virus surface antigen (HBsAg) |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests the selection of the equipment should take into account the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus (HBV) surface antigen (HBsAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay for hepatitis B virus (HBV) surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B surface antigen, HBsAg, hepatitis, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, CLEIA, HBV
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen for acute and chronic hepatitis B virus (HBV) infection: children, adolescents and adults
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 96%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.1 IU/mL
18 | Rate of invalid, error and unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 13%, Inter-assay precision (reproducibility): CV ≤ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay
22 | Specimen(s) stability | At 2–8 °C for 48 h and at –20 °C or lower for > 48 h
23 | Specimen(s) volume | ≤ 150 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 60 min
26 | End-point stability | Not applicable
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory: Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory: Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126 )
29 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirements once the reagents are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h and at –20 °C or lower for > 48 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h and at –20 °C or lower for > 48 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h and at –20 °C or lower for > 48 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirements once the reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 9 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Depends on the brand or model and whether it is bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the brand or model and whether it is bench-top or floor-standing
47 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideally, an uninterruptible power source
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: ≥ 2000 m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
53 | Displayed parameters | Digital display of test results, electronic and printed readout
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programmed database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
57 | Open or closed system | Both open and closed system
58 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
59 | Water requirements | For analysers with a wet system, deionized or distilled water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are contained in the test kit component, other reagents that the manufacturer specifies are required. Please see the package insert or instruction for use of the test kit.
64 | Reagent/kit size (number of tests) | ≥ 40 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other reagents and consumables specified by the manufacturer, disposable gloves and biohazard bags; and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste are required. Please see the package insert or instructions for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified in the kits/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spare parts that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. It should contain the actual results obtained during final quality control for release of each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply documentation for each method, providing the following detailed, comprehensive information: (i) health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal of the product; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More specific details will usually be found in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification; e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installating and initial setting up of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | Manufacturers must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests processed by the instrument, which will be dictated by the workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs, and training. Possibility of leasing an instrument for smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
82 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
83 | Estimated lifespan | 5 years after the start of use or installation on the condition of the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures performed under normal use and operating conditions
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. There is a WHO prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 5
Quantitative HBV nucleic acid test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | HBV nucleic acid test
3 | Specific type or variation (optional) | Quantitative test
4 | Alternative name/s (optional) | Quantitative HBV nucleic acid test
5 | Keywords (optional) | Hepatitis B virus, nucleic acid test, NAT, nucleic acid amplification test, NAAT, quantitative test, real-time polymerase chain reaction, PCR, transcription meditated amplification, TMA, HBV
Intended use
6 | Detection target | Hepatitis B virus nucleic acid
7 | Test purpose | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Nucleic acid test (manual method)
10 | Specimen type | Serum, plasma
11 | Testing population | Adults and children (0–18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | > 99%
15 | Clinical specificity | > 97%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Linear range (new addition) | 20–100 000 000 IU/mL (1.0–8.0 log10 IU/mL)
18 | Limit of detection | ≤ 20 IU/mL
19 | Rate of invalid, error and unreturnable results | < 5%
20 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV 9%; inter-assay precision (reproducibility): CV ≤ 10%
21 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
22 | Principle of the assay | Real-time polymerase chain reaction (PCR) targeting amplification of specific regions of HBV nucleic acid. The PCR product should be detected with fluorescent oligonucleotide probes that bind specifically to HBV nucleic acid. The test should be able to detect HBV genotypes A, B, C, D, E, F, G and H.
23 | Specimen(s) stability | At 2–8 °C for 24 h, at –20 °C or more for > 24 h
24 | Specimen(s) volume | 200–1000 µL
25 | Type of result | Quantitative
26 | Time to result | Within 4 h
27 | End-point stability | Not applicable
28 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 48 tests per 8 h per day
30 | Test limitations | The manufacturer should state any limitations of the test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
31 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify whether the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
32 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Transport stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected from humidity and direct sunlight
34 | Storage stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected from humidity and direct sunlight
35 | On-board stability of kit/reagents (temperature and humidity) | At the temperature inside the instrument
36 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C. Specific requirements once the reagents are opened should be clearly stated.
37 | Transport stability for specimen (temperature and humidity) | At 2–8 °C up to 1 day At –20 °C or colder for longer storage
38 | Storage stability for specimen (temperature and humidity) | At 2–8 °C up to 1 day At –20 °C or colder for longer storage
39 | In-use stability for specimen (temperature and humidity) | At 2–8 °C up to 1 day At –20 °C or colder for longer storage
40 | Transport stability for controls/calibrators (temperature and humidity) | At –15 °C to –30 °C. Repeated thawing and freezing should be avoided.
41 | Storage stability for controls/calibrators (temperature and humidity) | At –15 °C to –30 °C. Repeated thawing and freezing should be avoided.
42 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
43 | In-use stability for controls/calibrators (temperature and humidity) | At 2– 8 °C up to 1 day At –20 °C or colder for longer storage or at 15–28 °C
44 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
45 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
46 | Size of device (height x width x depth) | Nucleic acid extraction/isolation/purification: not applicable. Any commercially available nucleic acid extraction kit approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real-time PCR instrument: Depends on the specific brand or model. Any commercially available real-time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
47 | Weight of the device (kg) | Real-time PCR instrument: Depends on the brand or model. Any commercially available real-time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
48 | Power requirements | Real-time PCR instrument: Depends on the brand or model. Any commercially available real-time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
49 | Time to battery charge | As above
50 | Battery duration | As above
51 | Alternative charging options | As above
52 | Operating conditions (temperature and humidity) | As above
53 | User interface | As above
54 | Displayed parameters | As above
55 | Display languages | As above
56 | Built-in memory storage capacity | As above
57 | Diagnostic connectivity | As above
58 | Open or closed system | Both closed and open system
59 | Multi-disease testing capability/availability of a test menu | Analysis of several analytes
Infrastructure requirements
60 | Water requirements | Molecular-grade water
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
64 | Kit component – test, reagents, consumables (if relevant) | Reagents required for real-time PCR, including internal control, primers, probes, DNA polymerase and package Insert/Instructions for use
65 | Reagent/kit size (number of tests) | ≥ 24 tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, a commercially available kit for nucleic acid extraction, isolation and purification approved by regulatory authorities; HBV control and calibrator kits if not provided; 96 well-plates; tubes of PCR, covers, seals or foil for sealing the plates, aerosol barrier pipette tips, containers/tubes for specimens, disposable powder-free gloves, sodium hypochlorite solution at the concentration stated in the instructions for use or package insert for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop microcentrifuge for tubes, a plate centrifuge, a vortex mixer, a plate-shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated instrument for nucleic acid extraction, isolation and purification, a real-time PCR instrument or thermal cycler (specified by the test kit) with software, calibrated adjustable micropipettes, a cooling block, a timer, an electronic scale with hood and a cover for calibrating the micropipettes (scale that can weigh ≤ 0.001 g (1 µL)), a micropipette stand and an autoclave are required. (The autoclave cannot be used for nucleic products due to possible explosion and diffusion of nucleic products in the laboratory.) If the room temperature exceeds the upper limit for the operating conditions, an air conditioner is required.
68 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
69 | Instrument operator manual | For real-time PCR instrument: Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
71 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
72 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
73 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details will usually be provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration
Training, installation and use
75 | Pre-installation requirements (if relevant) | For real-time PCR instrument, access, space, water, power, waste, refrigeration and room temperature range should be assessed.
76 | Requirements for installation and calibration | For real-time PCR instrument: The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | For real-time PCR instrument, manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
78 | Warranty | For real-time PCR instrument, the manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
79 | Preventive maintenance | For real-time PCR instrument, the minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or according to the number of tests processed, which will be dictated by the workload.
80 | Corrective maintenance | For real-time PCR instrument, specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | For real-time PCR instrument, must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible leasing of small instruments
82 | Availability of spare parts post-warranty | For real-time PCR instrument, spare parts should be available for the life of the system, which should be a minimum of 5 years.
83 | Availability of in-country technical support | For real-time PCR instrument, required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
84 | Estimated lifespan | For real-time PCR instrument, in general, 5 years after the start of use or installation on condition of regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
85 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. See In vitro diagnostic medical devices used for the quantitative detection of hepatitis B DNA (Technical specifications series for submission to WHO prequalification – diagnostic assessment,TSS15) (https://extranet.who.int/pqweb/sites/default/files/documents/TSS-15_Hepatitis-B_NAT.pdf). No products for quantitative HBV nucleic acid testing were listed in the WHO list of prequalificated in vitro diagnostic products as of March 2023. See WHO public reports for in-vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 6 (Semi auto)
Quantitative HBV nucleic acid test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | HBV nucleic acid test
3 | Specific type or variation (optional) | Quantitative test
4 | Alternative name/s (optional) | Quantitative HBV nucleic acid test
5 | Keywords (optional) | Hepatitis B virus, nucleic acid test, NAT, nucleic acid amplification test, NAAT, quantitative test, real-time polymerase chain reaction, PCR, transcription meditated amplification, TMA, HBV
Intended use
6 | Detection target | Hepatitis B virus nucleic acid
7 | Test purpose | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Nucleic acid test (semi-automated method)
10 | Specimen type | Plasma and /or serum
11 | Testing population | Adults and children (0–18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | > 99% or sensitivity/limit of detection as analytical performance: ≤ 16 IU/mL
15 | Clinical specificity | > 95%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Linear range | 20–100 000 000 IU/mL (1.0–8.0 log10 IU/mL)
18 | Limit of detection | ≤ 20 IU/mL
19 | Rate of invalid, error and unreturnable results | < 5%
20 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV < 9%, inter-assay precision (reproducibility): CV < 11%
21 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
22 | Principle of the assay | Real-time polymerase chain reaction (PCR) for amplification of specific regions of HBV nucleic acid. The PCR product should be detected by fluorescent oligonucleotide probes that bind specifically to HBV nucleic acid. The test should be able to detect HBV genotypes A, B, C, D, E, F, G and H.
23 | Specimen(s) stability | At 2–8 °C for 48 h, at –20 °C or more for > 48 h
24 | Specimen(s) volume 200–1000 µL | 200–1000 µL
25 | Type of result | Quantitative
26 | Time to result | ≤ 4 h
27 | End-point stability | Not applicable
28 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 96 specimens per 8-h day
30 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
31 | Internal quality control | Internal control should be provided for sample preparation (nucleic acid extraction) and the subsequent real-time PCR to detect sample-specific reaction inhibition and failure to identify the reagent for consideration in interpreting the results. Re-testing should be recommended in case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
32 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Transport stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8°C protected from humidity and direct sunlight
34 | Storage stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8°C protected from humidity and direct sunlight
35 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's measuring unit
36 | In-use stability of kit/reagents (temperature and humidity) | At 15–30 ºC for ≤ 1day At 2–8 ºC for ≤ 3 days At –20 ºC or colder for longer storage
37 | Transport stability for specimen (temperature and humidity) | At 15–30 ºC for ≤ 1day At 2–8 ºC for ≤ 3 days At –20 ºC or colder for longer storage
38 | Storage stability for specimen (temperature and humidity) | At 15–30 ºC for ≤ 1day At 2–8 ºC for ≤ 3 days At –20 ºC or colder for longer storage
39 | In-use stability for specimen (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided.
40 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided
41 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided
42 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyser
43 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 ºC or at 15–28 ºC
44 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months
45 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
46 | Size of device (height x width x depth) | Automated instrument for nucleic acid extraction, isolation and purification: Depends on the brand or model. Any commercially available automated instruments for nucleic acid amplification and detection approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real-time PCR instrument: Depends on the brand or model. Any commercially available real-time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
47 | Weight of the device (kg), as above | As above
48 | Power requirements | Automated instrument for nucleic acid extraction, isolation and purification: 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz. Real-time PCR instrument: 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
49 | Time to charge battery | Not applicable and as above
50 | Battery duration | Not applicable and as above
51 | Alternative charging options | Ideally, an uninterruptible power source
52 | Operating conditions (temperature and humidity) | Automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Operating temperature range: 15–30 °C Relative humidity: ≤ 80% Altitude: ≤ 2000 m
53 | User interface | Automated instrument for nucleic acid extraction,isolation and purification and real-time PCR instrument: Simple test menu, simple PC as provided by manufacturer or built-in PC, option for connection to laboratory computer and middle ware
54 | Displayed parameters | Automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: digital display of test results, electronic and printed readout
55 | Display languages | English or local language
56 | Built-in memory storage capacity | Automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Two types of memory storage for instruments: those that contain the computer ring system (computer built in) and those that do not (computer not built in), depending on the brand or model. Please ask the manufacturer.
57 | Diagnostic connectivity | Automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases.
58 | Open or closed system | Automated instrument for nucleic acid amplification and detection and real-time PCR instrument: closed or open system
59 | Multi-disease testing capability/availability of a test menu | Real-time PCR instrument: multiple analytes
Infrastructure requirements
60 | Water requirements | Molecular-grade water
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
64 | Kit components – test, reagents, consumables (if relevant) | Reagents required for real-time PCR, including internal control, primers, probes, DNA polymerase and package insert or instructions for use
65 | Reagent/kit size (number of tests) | ≥ 24 tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection: alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, consumables depend on the kits. In general, the controls, calibrators, reaction tubes, specimen diluent, lysis reagent and buffer solution specified the kits are required. Please see the package insert or instructions for use. Common consumables for testing: sterile filter pipette tips, sterile microtubes, tube racks and disposable powder-free gloves, sodium hypochlorite solution at the concentration required by the instructions for use or package insert for sterilization before disposal of infectious waste. 70% ethanol and biohazard bags are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop microcentrifuge for tubes, a plate centrifuge, a vortex mixer, a plate-shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated nucleic acid extraction, isolation and purification instrument, a real-time PCR instrument or thermal cycler (specified by the test kit) with software, calibrated adjustable micropipettes, a cooling block, a timer, an electronic scale with hood and a cover for calibrating the micropipettes (scale can weigh ≤ 0.001 g (1 µL)), a micropipette stand and an autoclave are required. (The autoclave cannot be used for nucleic products because of possible explosion and diffusion of nucleic products in the laboratory.) If the room temperature is above the upper limit of the temperature for operating conditions, an air conditioner is required.
68 | Spare parts (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
69 | Instrument operator manual | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates that the lot was used.
71 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
72 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
73 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration
Training, installation and use
75 | Pre-installation requirements (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCT instrument: access, space, water, power, waste, refrigeration and room temperature range to be assessed
76 | Requirements for installation and calibration | For real-time PCR instrument: The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Manufacturers should provide detailed training in use and maintenance of the analyser before handing over system. Continuing training should be offered if appropriate.
Warranty and maintenance
78 | Warranty | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: The manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
79 | Preventive maintenance | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests processed, which will be dictated by the workload.
80 | Corrective maintenance | For real-time PCR instrument, specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCT instrument: Must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible leasing of small instruments
82 | Spare parts availability post-warranty | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCT instrument: Spare parts should be available for the life of the system, which should be a minimum of 5 years.
83 | In-country technical support availability | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCT instrument: Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
84 | Estimated lifespan | For automated instrument for nucleic acid amplification and detection and real-time PCR instrument: 5 years after the start of use/installation on condition of regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
85 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. See In vitro diagnostic medical devices used for the quantitative detection of Hepatitis B DNA (Technical specifications series for submission to WHO prequalification – diagnostic assessment,TSS15) (https://extranet.who.int/pqweb/sites/default/files/documents/TSS-15_Hepatitis-B_NAT.pdf). No products for quantitative HBV nucleic acid tests were listed in the WHO list of prequalified vitro diagnostic products as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 7 (Fully auto)
Quantitative HBV nucleic acid test |  | This technical specification is for a specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests the selection of the equipment should also take into account other tests performed in the facility and not just the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | HBV nucleic acid test
3 | Specific type or variation (optional) | Quantitative test
4 | Alternative name/s (optional) | Quantitative HBV nucleic acid test
5 | Keywords (optional) | Hepatitis B virus, nucleic acid test, NAT, nucleic acid amplification test, NAAT, quantitative test, real-time polymerase chain reaction, PCR, transcription meditated amplification, TMA, HBV
Intended use
6 | Detection target | Hepatitis B virus nucleic acid
7 | Test purpose | To stage chronic HBV infection, to determine the need for treatment (including use of antivirals in mothers to prevent mother-to-child transmission) and to monitor response to treatment
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Nucleic acid test (automated method)
10 | Specimen type | Serum, plasma
11 | Testing population | Adults and children (0–18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | Clinical sensitivity as clinical performance: No data Sensitivity/limit of detection as analytical performance: < 10 IU/mL
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Linear range (new addition): 10 IU/mL to 100 000 000 IU/mL | 20–100 000 000 IU/mL (1.0–8.0 log10 IU/mL)
18 | Limit of detection | ≤ 20 IU/mL
19 | Rate of invalid, error and unreturnable results | Not applicable
20 | Precision (repeatability/reproducibility) | When positive samples with not very low concentration are used: Intra-assay precision (repeatability): CV ≤ 20% Inter-assay precision (reproducibility):CV ≤ 20%
21 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
22 | Principle of the assay | Real-time polymerase chain reaction (PCR) or transcription meditated amplification (TMA) for amplification of specific regions of HBV nucleic acid. The PCR or TMA product should be detected by fluorescent oligonucleotide probes that bind specifically to HBV nucleic acid. A real-time PCR or TMA machine should be used to detect the fluorescence. The test should be able to detect HBV genotypes A, B, C, D, E, F, G and H.
23 | Specimen(s) stability | Whole blood at 2–30 ºC for ≥ 24 h Plasma and serum at 2–8 ºC for ≥ 2 days Plasma and serum at –18 ºC to –20 ºC for ≤ 6 weeks Plasma and serum at –60 ºC or colder for ≥ 6 weeks
24 | Specimen(s) volume | 400–1200 µL
25 | Type of result | Quantitative
26 | Time to result | 1–3.5 h
27 | End-point stability | Not applicable
28 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | For low sample volume laboratories: 6–24 tests per 8-h day For middle sample volume laboratories: 96–400 tests per 8-h day For high sample volume laboratories: > 400 tests per 8-h day
30 | Test limitations | The manufacturer should state any limitations of the test, inclusing conditions that have not been validated with the test, e.g. testing population, sample types, users.
31 | Internal quality control | Internal control should be provided for sample preparation (nucleic acid extraction) and for the subsequent real-time reverse transcription PCR or TMA to detect sample-specific reaction inhibition and systemic malfunction due to failure to identify the reagent, and be considered in interpreting the results. Re-testing should be recommended in the procedure in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
32 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C protected from humidity and direct sunlight in an upright position Ideally, > 8 °C
34 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C protected from humidity and direct sunlight in an upright position Ideally, > 8 °C
35 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyser's reagent unit or reagent storage
36 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C protected from humidity and direct sunlight in an upright position Ideally, > 8 °C
37 | Transport stability for specimen (temperature and humidity) | Whole blood at 2–30 ºC for ≥ 24h Plasma and serum at 2–8 ºC for ≥ 3 days Plasma and serum at –18 ºC to –20 ºC for ≤ 6 weeks Plasma and serum at –60 ºC or colder for ≥ 6 weeks
38 | Storage stability for specimen (temperature and humidity) | Whole blood at 2–30 ºC for ≥ 24h Plasma and serum at 2–8 ºC for ≥ 3 days Plasma and serum at –18 ºC to –20 ºC for ≤ 6 weeks Plasma and serum at –60 ºC or colder for ≥ 6 weeks
39 | In-use stability for specimen (temperature and humidity) | Whole blood at 2–30 ºC for ≥ 24h Plasma and serum at 2–8 ºC for ≥ 3 days Plasma and serum at –18 ºC to –20 ºC for ≤ 6 weeks Plasma and serum at –60 ºC or colder for ≥ 6 weeks
40 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided.
41 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25°C Ideally, > 8 °C The freeze–thaw cycle depends on the test kits and systems.
42 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyser
43 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 ºC or at 15–25 ºC Ideally, > 8 °C
44 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture of kits
45 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
46 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
47 | Weight of the device (kg), as above | Depends on the brand or model and whether bench-top or floor-standing
48 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
49 | Time to charge battery | Not applicable
50 | Battery duration | Not applicable
51 | Alternative charging options | Ideally, uninterruptible power source
52 | Operating conditions (temperature and humidity) | Operating temperature range: ≥ 18–28 °C Humidity: ≥ 30–65%
53 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
54 | Displayed parameters | Digital display of test results, electronic and printed readout
55 | Display languages | English or local language
56 | Built-in memory storage capacity | Two types of memory storage: analysers that contain the computer ring system (computer built-in) and those that do not (computer not built-in), depending on the brand or model. Please ask the manufacturer.
57 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases.
58 | Open or closed system | Both closed and semi-closed
59 | Multi-disease testing capability/availability of a test menu | Analysis of several analytes
Infrastructure requirements
60 | Water requirements | No
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
64 | Kit component – test, reagents, consumables (if relevant) | Reagents required for PCR or TMA (reagents for DNA extraction and amplification) or Test trays or test cartridges containing the reagents required for PCR or TMA (reagents for DNA extraction and amplification) and package inserts or instructions for use.
65 | Reagent/kit size (number of tests) | ≥ 10 tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, consumables depend on the kits. In general, controls, calibrators, reaction tubes, specimen diluent, lysis reagent and buffer solution specified for the kits are required. Please see the package inserts or instructions for use. Common consumables for testing (sterile filter pipette tips, sterile microtubes, tube racks and disposable powder-free gloves, sodium hypochlorite solution at the concentration required by the instructions for use/package insert for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags) are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum and plasma, a centrifuge, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, an analyser and specification file or software for the analyser are required. A biological safety cabinet approved for working with infectious materials is desirable. Common equipment (such as a timer, a vortex mixer, a centrifuge, calibrated micropipettes (10–1000 μL), a micropipette stand, an electronic scale with hood and a cover for calibrating the micropipettes (the scale can weigh ≤ 0.001 g (1 µL)), a refrigerator, a freezer and an autoclave) for testing are required. (The autoclave should not be used for nucleic products because of possible scattering of nucleic products). If the room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
68 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
69 | Instrument operator manual | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also provide recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
72 | Material safety data sheet | Required. The manufacturer should supply detailed, documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
73 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g. health, physical or environmental . More details are usually provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration
Training, installation and use
75 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
76 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
Warranty and maintenance
78 | Warranty | Manufacturers must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
79 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will be dictated by the workload.
80 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing an instrument for smaller devices.
82 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 7 years.
83 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
84 | Estimated lifespan | 7 years after the start of use or installation on the condition of the regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
85 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. See In vitro diagnostic medical devices used for the quantitative detection of Hepatitis B DNA (Technical specifications series for submission to WHO prequalification – diagnostic assessment,TSS15) (https://extranet.who.int/pqweb/sites/default/files/documents/TSS-15_Hepatitis-B_NAT.pdf). No quantitative HBV nucleic acid tests were listed on the WHO list of prequalificated in vitro diagnostic products as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 8
Hepatitis B e antigen (HBeAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
This assay is within the scope of WHO Prequalification as of March 2023. See In vitro diagnostic medical devices used for the quantitative detection of Hepatitis B DNA (Technical specifications series for submission to WHO prequalification – diagnostic assessment,TSS15) (https://extranet.who.int/pqweb/sites/default/files/documents/TSS-15_Hepatitis-B_NAT.pdf). No quantitative HBV nucleic acid tests were listed on the WHO list of prequalificated in vitro diagnostic products as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr). | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus e antigen (HBeAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for hepatitis B e antigen (HBeAg)
5 | Keywords (optional) | Hepatitis B virus e antigen, HBeAg, hepatitis, microtitre plate enzyme immunoassay, HBV, quantitative
Intended use
6 | Detection target | Hepatitis B virus e antigen (HBeAg)
7 | Test purpose | To stage chronic HBV infection and to determine the need for treatment
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Adults and children (0–18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 99%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.2 PEI U/mL (PEI: Paul Ehrlich Institut)
18 | Rate of invalid, error and unreturnable results | No information
19 | Precision (repeatability/reproducibility) | Intra-assay precision (reproducibility): CV ≤ 11% Inter-assay precision (reproducibility at several sites): CV ≤ 18%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Enzyme immunoassay
22 | Specimen(s) stability | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
23 | Specimen(s) volume | ≤ 100 µL
24 | Type of result | Qualitative or qualitative
25 | Time to result | 180–210 min
26 | End-point stability | 30 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify whether the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once controls are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | 11 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay and enzyme-linked immunosorbent assay. See the instructions for use or the package insert of the kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit component – test, reagents, consumables (if relevant) | Coated wells, strips with coated wells, positive and negative controls, conjugate, conjugate diluent, wash buffer, chromogen substrate, stop solution and package insert or instructions for use
64 | Reagent/kit size (number of tests) | ≥ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader, enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer, enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Docymentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually contained in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacturer was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 9 (Auto)
Hepatitis B virus e antigen (HBeAg) |  | This technical specification is for a specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, selection of the equipment should also take into account other tests performed in the facility and not just the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B e antigen (HBeAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay for hepatitis B e antigen (HBeAg)
5 | Keywords (optional) | Hepatitis B virus e antigen, HBeAg, hepatitis, chemiluminescent immunoassay, chemiluminescent enzyme immunoassay, HBV
Intended use
6 | Detection target | Hepatitis B virus e antigen (HBeAg),
7 | Test purpose | To stage chronic HBV infection and to determine the need for treatment
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (automated method)
10 | Specimen type | Plasma and serum or serum
11 | Testing population | Adults and children (0–18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 97%
15 | Clinical specificity | ≥ 98%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.5 PEI U/mL(PEI: Paul Ehrlich Institut)
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 10%, inter-assay precision (reproducibility): CV ≤ 12%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay
22 | Specimen(s) stability | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
23 | Specimen(s) volume | ≤ 150 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 60 min
26 | End-point stability | Not applicable
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory: Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126 )
29 | Test limitations | The manufacturer should state any test limitations, including the conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify whether the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement for controls or calibrators once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | 11 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
47 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideally, an uninterruptible power source
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: 2000 m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
53 | Displayed parameters | Digital display of test results, electronic and printed readout
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
57 | Open or closed system | Open and closed systems
58 | Multi-disease testing capabiliy/availability of a test menu | Multi-disease testing
Infrastructure requirements
59 | Water requirements | For analysers with a wet assay system, deionized water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry assay system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are contained in the test kit, other reagents specified by the manufacturer are required. Please see the package insert or instructions for use of the test kit.
64 | Reagent/kit size (number of tests) | ≥ 20 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables specified by the manufacturer, disposable gloves and biohazard bags and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste are required. Please see the package insert or instructions for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified for the kit/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually available in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | Manufacturers must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of the warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which is dictated by the workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing instruments for smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
82 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
83 | Estimated lifespan | 5 years after the start of use or installation on the condition that the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 10
IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay for IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc)
5 | Keywords (optional) | IgM antibodies to hepatitis B core antigen, hepatitis B virus, microtitre plate enzyme immunoassay, HBV
Intended use
6 | Detection target | IgM-specific antibodies to hepatitis B core antigen
7 | Test purpose | To aid in the diagnosis of acute HBV infection in the context of an outbreak investigation
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 96%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | 50 PEI unit/mL (PEI: Paul Ehrlich Institute)
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 13%, inter-assay precision (reproducibility): CV ≤ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Enzyme immunoassay, microwells coated with anti-human IgM antibody
22 | Specimen(s) stability | At 2–8 °C for 72 h and at –20 °C or lower for > 72 h
23 | Specimen(s) volume | 10–100 µL
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | 120–180 min
26 | End-point stability | 10–20 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify whether the performance of the assay is as qualified. Re-testing should be recommended in cases of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirements once the reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C beyond 72 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for 12 months
38 | In-use stability for specimen collection (temperature and humidity) | At –20 °C for 12 months. Avoid repeated freezing and thawing (maximum twice).
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once controls are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay or enzyme-linked immunosorbent assay. See the instruction for use or the package insert of the kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit component – test, reagents, consumables (if relevant) | Coated plate with wells, conjugate and positive and negative control sera, specimen diluent, conjugate, conjugate diluent, wash buffer, substrate concentrate, substrate diluent, stop solution and package insert/instructions for use.
64 | Reagent/kit size (number of tests) | ≥ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay, spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with a hood and a cover for calibrating the micropipettes that can weigh ≥ 0.001 g (1 µL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This is usually contained in material safety data sheet.
73 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 11 (Auto)
IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) |  | This technical specification is for a specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests the selection of the equipment should also take into account other tests performed in the facility and not just the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay for IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc) or chemiluminescent enzyme immunoassay for IgM-specific antibodies to hepatitis B core antigen (IgM anti-HBc)
5 | Keywords (optional) | IgM antibodies to hepatitis B core antigen, hepatitis B virus, chemiluminescent immunoassay, CIA, chemiluminescent enzyme immunoassay, CLIA, HBV
Intended use
6 | Detection target | IgM-specific antibodies to hepatitis B core antigen
7 | Test purpose | To aid in the diagnosis of acute HBV infection in the context of an outbreak investigation
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 96%
15 | Clinical specificity | ≥ 98%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 23 PEI U/mL (PEI: Paul Ehrlich Institute)
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 10%; inter-assay precision (reproducibility): CV ≤ 21%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay
22 | Specimen(s) stability | At 2–8 °C for 24 h and at –20 °C beyond 72 h
23 | Specimen(s) volume | ≤ 170 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 50 min
26 | End-point stability | Not applicable
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory Specimen per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126 )
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify whether the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 72 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 72 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 72 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2– 8 °C, protected from all heat and light sources
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2– 8 °C, protected from all heat and light sources
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement once controls or calibrators are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
47 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideal: Uninterruptible power source required
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C, humidity: < 80%, altitude: ≤ 2000 m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
53 | Displayed parameters | Digital display of test results, electronic and printed readout
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyser for low sample volume laboratory : ≥ 800 tests Analyser for middle sample volume laboratory : ≥ 2000 tests Analyser for high sample volume laboratory : ≥ 10 000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
57 | Open or closed system | Open and closed systems
58 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
59 | Water requirements | For analysers with a wet assay system, deionized water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry assay system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit component – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are contained in the test kit component, other reagents specified by the manufacturer are required. Please see the package insert or instructions for use of the test kit.
64 | Reagent/kit size (number of tests) | ≥ 42 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables specified by the manufacturer, disposable gloves and biohazard bags and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, are required. Please see the package insert or instructions for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified by the kits/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential sources of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More specific details are usually contained in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | Manufacturers must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which is dictated by the workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing instruments for smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
82 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical reqirements of the user.
Decommissioning
83 | Estimated lifespan | 5 years after the start of use and installation on the condition that the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 12
Antibodies to hepatitis B virus surface antigen (anti-HBs)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Antibodies to hepatitis B virus surface antigen (anti-HBs) immunoassay
3 | Specific type or variation (optional) | Quantitative and/or qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for antibodies to hepatitis B virus surface antigen (anti-HBs)
5 | Keywords (optional) | Antibodies to hepatitis B virus surface antigen, anti-HBs, hepatitis, microtitre plate enzyme immunoassay, quantitative, HBV
Intended use
6 | Detection target | Antibodies to hepatitis B virus surface antigen (anti-HBs)
7 | Test purpose | To determine immune status after HBV vaccination
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance chacateristics
14 | Clinical sensitivity | > 99%
15 | Clinical specificity | > 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 10 U/mL
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV 13%; inter-assay precision (repeatability): CV 13%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Enzyme immunoassay, microwells coated with antibodies to hepatitis B virus surface antigen (anti-HBs)
22 | Specimen(s) stability | At 2–8 °C for 72 h and at –20 °C for longer storage. Avoid successive freezing and thawing.
23 | Specimen(s) volume | ≤ 100 µL
24 | Type of result | Qualitative (reactive or non-reactive)
25 | Time to result | 180–240 min
26 | End-point stability | ≤ 30 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for longer storage
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for longer storage
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for longer storage
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | 12 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multidisease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay or enzyme-linked immunosorbent assay. See the instructions for use or package insert of the kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | Coated plate with wells, positive and negative control sera, specimen diluent, conjugate, conjugate diluent, wash buffer, substrate concentrate, substrate diluent, stop solution and package insert or instructions for use.
64 | Reagent/kit size (number of tests) | ≧ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001g (1 µL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Document size
68 | Instrument operator manual | Not applicable
69 | Instructions for use | The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing cente with the dates on which the lot was used.
71 | Material safety data sheet | Hepatitis B virus (HBV) infection
Environmental and safety requirements
72 | Hazard classification | Potential sources of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This is usually stated in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.
Note on D21: Elisabeth Heseltine: Infants are < 12 months

## 13 (Auto)
Antibodies to hepatitis B virus surface antigen (anti-HBs) |  | This technical specification is for a specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests the selection of the equipment should also take into account other tests performed in the facility and not just the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Antibodies to hepatitis B virus surface antigen (anti-HBs) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay for antibodies to hepatitis B virus surface antigen (anti-HBs) or chemiluminescent enzyme immunoassay for antibodies to hepatitis B virus surface antigen (anti-HBs)
5 | Keywords (optional) | Antibodies to hepatitis B virus surface antigen, anti-HBs, hepatitis, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, CLEIA, HBV
Intended use
6 | Detection target | Antibodies to hepatitis B virus surface antigen (anti-HBs)
7 | Test purpose | To determine immune status after HBV vaccination
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 95%
15 | Clinical specificity | ≥ 96%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 10 U/mL
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 13%; inter-assay precision (reproducibility): CV ≤ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Chemiluminescentimmunoassay or chemiluminescent enzyme immunoassay
22 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 24 h or at –20 °C for longer storage.
23 | Specimen(s) volume | ≤ 150 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 60 min
26 | End-point stability | Not applicable
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory: Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62) Middle sample volume laboratory: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory: Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126)
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyser's reagent unit or reagent storage
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C or lower for > 24 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources. Any specific requirement once controls or calibrators are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≤ 10 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when the product arrives at the point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
47 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideally, an uninterruptible power source
51 | Operating conditions (temperature, humidity and altitude) | Operating temperature range: 18–30 °C; humidity: < 80%; altitude: ≤ 2000 m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
53 | Displayed parameters | Digital display of test results, electronic and printed readout
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for mddle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
57 | Open or closed system | Both open and closed systems
58 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
59 | Water requirements | For analysers with a wet assay system, deionized water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry assay system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are contained in the test kit component, other reagents specified by the manufacturer are required. Please see the package insert or instruction for use of the test kit.
64 | Reagent/kit size (number of tests) | ≥ 20 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other reagents and consumables specified by the manufacturer, disposable gloves, biohazard bags and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste are required. Please see the package insert or instructions for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified for the kits/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also state the recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling, storage, use and disposal of the product; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential sources of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually contained in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | The manufacturer must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will be dictated by the workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing for smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
82 | Availability of in-country technical support | Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
83 | Estimated lifespan | 5 years after the start of use or installation on the condition that the regular maintenanceand inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was not within the scope of WHO Prequalification as of March 2023. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 14
Hepatitis B virus surface antigen (HBsAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus surface antigen (HBsAg)
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for hepatitis B virus surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B virus surface antigen, HBsAg, hepatitis, screening, rapid diagnostic test
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen donated blood for HBV
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
11 | Testing population | Adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | For blood-screening laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 97%
15 | Clinical specificity | ≥ 98%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 2 IU/mL
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Lateral flow assay
22 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. Specimens can be stored at 2–8 °C. Do not freeze whole-blood specimens.
23 | Specimen(s) volume | 25–100 µL
24 | Type of result | Qualitative (reactive or non-reactive)
25 | Time to result | 15–30 min
26 | End-point stability | 5–10 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Maximum 2 steps
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 2 specimens per hour per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | A built-in quality control line that must appear when a specimen is added is part of internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 4–30 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature or use of a centrifuge will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | Test cards, test cassettes or test strips
64 | Reagent/kit size (number of tests) | ≥ 20 tests
65 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipettes are not provided, tips for micropipettes are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettes for venous whole blood are not provided, tips for micropipettes are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified for the kit/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually provided in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacturer was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay in RDT format was within the scope of WHO Prequalification as of 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 15
Hepatitis B virus surface antigen (HBsAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus surface antigen (HBsAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for hepatitis B virus surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B surface antigen, HBsAg, hepatitis, microtitre plate enzyme immunoassay, HBV
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen donated blood for HBV
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | For blood-screening laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 99%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.2 IU/mL
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≥ 11%; inter-assay precision (reproducibility): CV ≥ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Enzyme immunoassay, microwells coated with antibodies to HBsAg
22 | Specimen(s) stability | At 2–8 °C for 72 h and at –20 °C beyond 72 h
23 | Specimen(s) volume | ≤ 100 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 240 min
26 | End-point stability | 10–20 min
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify the performance of the assay as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at at –20 °C beyond 72 h
37 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at at –20 °C beyond 72 h
38 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at at –20 °C beyond 72 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once controls and reagents are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay or enzyme-linked immunosorbent assay. See the instructions for use or package insert of kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components – test, reagents, consumables (if relevant) | Coated plate with wells, conjugate and HBsAg-positive and -negative control sera and package insert or instructions for use
64 | Reagent/kit size (number of tests) | ≥ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a plate cover, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre, with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | Potential sources of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually provided in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 16(Auto)
Hepatitis B virus surface antigen (HBsAg) |  | This technical specification is for a specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests the selection of the equipment should also take into account other tests performed in the facility and not just the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-31
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | Generic name | Hepatitis B virus (HBV) surface antigen (HBsAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay for hepatitis B virus (HBV) surface antigen (HBsAg)
5 | Keywords (optional) | Hepatitis B surface antigen, HBsAg, hepatitis, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, CLEIA, HBV
Intended use
6 | Detection target | Hepatitis B virus surface antigen (HBsAg)
7 | Test purpose | To screen donated blood for HBV
8 | Specific disorder or condition | Hepatitis B virus (HBV) infection
9 | Assay format | Immunoassay (automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
13 | Level of the health-care system | For blood-screening laboratories
Performance characteristics
14 | Clinical sensitivity | ≥ 99%
15 | Clinical specificity | ≥ 99%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HBV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
17 | Limit of detection | ≤ 0.1 IU/mL
18 | Rate of invalid, error, unreturnable results | < 5%
19 | Precision (repeatability/reproducibility) | When positive samples are used, intra-assay precision (repeatability): CV ≤ 13%; inter-assay precision (reproducibility): CV ≤ 15%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay
22 | Specimen(s) stability | At 2–8 °C for 48 h and at –20 °C or lower for > 48 h
23 | Specimen(s) volume | ≤ 150 µL
24 | Type of result | Qualitative
25 | Time to result | ≤ 60 min
26 | End-point stability | Not applicable
27 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory: Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62) Middle sample volume laboratory: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory: Specimen per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126 )
29 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
30 | Internal quality control | Positive and negative controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
31 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
33 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyser's reagent unit or reagent storage
35 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At –8 °C for 48 h and at –20 °C or lower for > 48 h
37 | Storage stability for specimen collection (temperature and humidity) | At –8 °C for 48 h and at –20 °C or lower for > 48 h
38 | In-use stability for specimen collection (temperature and humidity) | At –8 °C for 48 h and at –20 °C or lower for > 48 h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement for controls or calibrators once they are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 9 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
47 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
48 | Time to charge battery | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideally, an uninterruptible power source
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: ≤ 2000 m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
53 | Displayed parameters | Digital display of test results, electronic and printed readout
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programmed database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
57 | Open or closed system | Both open and closed systems
58 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
59 | Water requirements | Deionized or distilled water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
63 | Kit components: test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are contained in the test kit, other reagents specified by the manufacturer are required. Please see the package insert or instructions for use of the test kit.
64 | Reagent/kit size (number of tests) | ≥ 40 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables specified by the manufacturer, disposable gloves, biohazard bags and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste are required. Please see the package insert or instructions for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified by the kit/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also state the recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
71 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling, storage, use and disposal of the product; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential sources of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More specific details are usually provided in the material safety data sheet.
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to the users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training in the use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | The manufacturer must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will be dictated by the workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures from information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing instruments or smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
82 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
83 | Estimated lifespan | 5 years after the start of use or installation, on the condition that the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## Acknowledgement and Purpose
Purpose and aim
The Model List of Essential In-vitro Diagnostics (EDL) is a list of recommended types of tests to support diagnosis of infectious diseases and noncommunicable diseases at each level of the health-care system. EDL serves as an evidence-based reference point for countries to develop their own national lists to guide how they choose and use IVDs. The EDL is also published in an electronic format (eEDL). Detailed technical specification are required to guide countries in procuring appropriate tests. Technical specifications for IVDs consist of criteria and baseline requirements to ensure good quality, safety, performance and efficacy. Technical specifications for IVDs listed in the EDL will be published in an electronic format in the eEDL. The purpose of technical specifications for EDL is to advance universal health coverage, address health emergencies and promote healthier populations by helping countries to procure appropriate IVD products. The technical specifications provide detailed technical information to support decisions on procurement of good-quality, safe IVDs to improve access to in vitro diagnostic testing in low- and middle-income countries in particular. These technical specifications can be used in various settings, from regular health facilities to humanitarian settings and for emergencies.
Acknowledgements
Technical contributions: Nikhil Moorchung (WHO consultant, India), Emma English (WHO consultant, United Kingdom), Jitendra Kumar (WHO consultant, India), Olga Perovic (WHO consultant, South Africa), Ehab Saad (WHO consultant, South Africa), Francis Gabriel Moussy (WHO headquarters), Ana Aceves Capri (WHO headquarters), Naofumi Hashimoto (WHO headquarters), Adriana Velazquez Berumen (WHO headquarters), Philippa Easterbrook (WHO headquarters), Niklas Luhmann (WHO headquarters) and Ute Ströher (WHO headquarters).
Conflicts of interest and declarations of interest
All the WHO consultants involved in preparing these technical specifications provided declarations of interests, and no conflicts were identified.
Funding
Funding and in-kind contributions to this project were provided by the following organizations: • The Bill & Melinda Gates Foundation • The Foreign, Commonwealth and Development Office, United Kingdom • The Economic Research Institute for the Association of Southeast Asian Nations and East Asia
Updating
The document will be updated regularly, at least every 3 years.

## Process
Process
Technical specifications define the minimum requirements for ensuring the quality, safety and efficacy of the products. Development of technical specifications consists of the six steps listed below.
•Step 1: Review existing WHO guidance and specifications for the procurement, donation or lease of good-quality IVDs to develop a common template of generic specifications. The guidance documents analysed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment, Laboratory diagnosis of group A streptococcal infections, WHO Model List of Essential In Vitro Diagnostics and WHO internal procurement templates.
•Step 2: WHO commissions a team of external consultants to develop technical specifications for the IVDs listed in the WHO-EDL. Development of specifications comprises: horizon scanning of existing tests and platforms; analysis of products on the market after approval by stringent regulatory authorities; analysis of international, regional and country standards; review of WHO guidance and guidelines for all the diseases included in the EDL; review of peer-reviewed publications, textbooks and published qualitative and quantitative comparative studies of tests and platforms; comparison of test specifications (i.e. accuracy, precision, positive and negative predictive value, likelihood ratios), comparison and review of the package inserts and user manuals of products approved by stringent regulatory authorities; comparison and review of documents and data on the safety and performance of approved products on the websites of stringent regulatory authorities; and the results of external quality assessments and technical expert opinion. Each template is drafted or reviewed by at least two qualified technical experts. All consultants provided declarations of Interest, and no conflicts were identified.
•Step 3: Internal review of all draft technical specifications by the EDL secretariat •Step 4: Review of draft technical specifications by relevant disease programmes •Step 5: 3-week public consultation on the draft technical specifications by posting them on the WHO website and sharing them on relevant networks and with non-State actors in official relations with WHO, such as the International Federation of Biomedical Laboratory Science. •Step 6: Compilation of all feedback and finalization of technical specifications before adding them to the electronic EDL.

## Evidence and Input
Evidence and Input
Sources of evidence:
1. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment. Geneva: World Health Organization; 2017 (https://apps.who.int/iris/handle/10665/255577). 2. Guidelines on hepatitis B and C testing. Geneva: World Health Organization; 2017 (https://www.who.int/publications/i/item/978924154998). 3. Guidelines for the prevention, care and treatment of persons with chronic hepatitis B infection. Geneva: World Health Organization; 2015 (https://apps.who.int/iris/handle/10665/154590). 4. Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS-13 – Rapid diagnostic tests to detect hepatitis B surface antigen. Geneva: World Health Organization; 2022 (https://apps.who.int/iris/bitstream/handle/10665/366331/9789240063068-eng.pdf?sequence=1&isAllowed=y). 5. Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS-14 – Immunoassays to detect hepatitis B virus surface antigen. Geneva: World Health Organization; 2023 (https://apps.who.int/iris/bitstream/handle/10665/366463/9789240063082-eng.pdf). 6. Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS-15 – In vitro diagnostic medical devices used for the quantitative detection of hepatitis B DNA. Geneva: World Health Organization; 2022 (https://apps.who.int/iris/bitstream/handle/10665/366330/9789240064393-eng.pdf?sequence=1&isAllowed=y).
7. Reissinger A, Volkers P, Scheiblauer H, Nick S, WHO Expert Committee on Biological Standardization et al. Collaborative study to establish a World Health Organization international standard for hepatitis B e antigen (‎HBeAg)‎. Geneva: World Health Organization; 2013 (https://apps.who.int/iris/handle/10665/96313). 8. Hepatitis B surface antigen assays: operational characteristics (phase I). Report 2. Geneva: World Health Organization; 2004 (https://www.paho.org/en/node/58369). 9. Hepatitis B surface antigen assays: operational characteristics (phase I). Report 1. Geneva: World Health Organization; 2001 (https://apps.who.int/iris/handle/10665/66924?locale-attribute=pt&).
10. Hepatitis B market report 2022. Highlighting the latest trends in HBV treatment and diagnostics. Boston (MA): Clinton Health Access Initiative, Inc., 2022 (https://www.clintonhealthaccess.org/wp-content/uploads/2022/12/CHAI-HBV-Market-Report-2022_vf.pdf). 11. HCV market intelligence report 2021 and preliminary HBV market insights. Boston (MA): Clinton Health Access Initiative, Inc.; 2021 (https://chai19.wpenginepowered.com/wp-content/uploads/2021/08/Hepatitis-C-Market-Report_2021-FINAL-1.pdf). 12. List of HIV diagnostic test kits and equipment classified according to the Global Fund quality assurance policy. Version 48, 29 March 2023. Geneva: The Global Fund to Fight AIDS, Malaria and Tuberculosis; 2023 (https://www.theglobalfund.org/media/5878/psm_productshiv-who_list_en.pdf).
13. Committee for Evaluation of In Vitro Diagnostic Devices. Re-evaluation of HBsAg detection kits approved for marketing in Japan. Jpn J Infect Dis. 2001;54(5):201–7. PMID: 11754163. 14. Stockdale AJ, Silungwe NM, Shawa IT, Kreuls B, Gordon MA, Geretti AM. Diagnostic performance evaluation of hepatitis B e antigen rapid diagnostic tests in Malawi. BMC Infect Dis. 2021;21(1):487. doi:10.1186/s12879-021-06134-3. 15. Seck A, Ndiaye F, Maylin S, Ndiaye B, Simon F, Funk AL et al. Poor sensitivity of commercial rapid diagnostic tests for hepatitis B e antigen in Senegal, West Africa. Am J Trop Med Hyg. 2018;99(2):428–34. doi:10.4269/ajtmh.18-0116. 16. Nick S, Scheiblauer H. Sensitivities of CE-marked HIV, HCV, and HBsAg assays. J Med Virol. 2007;79:S59–64. doi:10.1002/jmv.20965. 17. Villar LM, Cruz HM, Barbosa JR,Bezerra CS, Portilho MM, de Paula Scalioni L. Update on hepatitis B and C virus diagnosis. World J Virol. 2015;4(4):323–42. doi:10.5501/wjv.v4.i4.323. 18. Lange B, Roberts T, Cohn J, Greenman J, Camp J, Ishizaki A et al. Diagnostic accuracy of detection and quantification of HBV-DNA and HCV-RNA using dried blood spot (DBS) samples – a systematic review and meta-analysis. BMC Infect Dis. 2017;17(Suppl 1):693. doi:10.1186/s12879-017-2776-z.
19. Amini A, Varsaneux O, Kelly H. Diagnostic accuracy of tests to detect hepatitis B surface antigen: a systematic review of the literature and meta-analysis. BMC Infect Dis. 2017;17(Suppl 1):698. doi:10.1186/s12879-017-2776-z. 20. Huzly D, Schenk T, Jilg W, Neumann-Haefelin D. Comparison of nine commercially available assays for quantification of antibody response to hepatitis B virus surface antigen. J Clin Microbiol. 2008;46(4):1298–306. doi:10.1128/JCM.02430-07. 21. Dabbs DJ, editor. Diagnostic immunohistochemistry: theranostic and genomic applications, 5th edition. Amsterdam: Elsevier; 2018 (https://www.eu.elsevierhealth.com/diagnostic-immunohistochemistry-9780323477321.html). 22. Cadamuro J, Lippi G, von Meyer A, Ibarz M, van Dongen E, Lases5 et al. European survey on preanalytical sample handling – Part 2: Practices of European laboratories on monitoring and processing haemolytic, icteric and lipemic samples. On behalf of the European Federation of Clinical Chemistry and Laboratory Medicine (EFLM) Working Group for the Preanalytical Phase (WG-PRE). Biochem Med (Zagreb). 2019;29(2):020705. doi:10.11613/BM.2019.020705. 23. Ogrodnik PJ. In vitro diagnostic devices: Class 1. Sterile and measuring devices. Amsterdam: Elsevier Science; 2014. ISBN-10. 0128000252.
24. Medical laboratories – requirements for quality and competence. ISO 15189:2012. Geneva: International Standards Organization; 2022 (https://www.iso.org/standard/56115.html). 25. FDA guidance for industry: sterile drug products produced by aseptic processing – current good manufacturing practice. College Park (MD): Department of Health and Human Services, Food and Drug Administration, Center for Drug Evaluation and Research; 2004 (https://www.gmp-compliance.org/guidelines/gmp-guideline/fda-guidance-for-industry-sterile-drug-products-produced-by-aseptic-processing-current-good-manufacturing-practice-september-2004).
26. Pisano DJ, Mantus DS, editors. FDA regulatory affairs: A guide for prescription drugs, medical devices, and biologics. Boca Raton (FL): CRC Press; 2008. doi:10.3109/9781420073553. 27. Fiedler BA, editor. Managing medical devices within a regulatory framework. Amsterdam: Elsevier Science; 2016 (https://www.elsevier.com/books/managing-medical-devices-within-a-regulatory-framework/fiedler/978-0-12-804179-6). 28. Tobin JJ, Walsh G. Medical product regulatory affairs: Pharmaceuticals, diagnostics, medical devices. Hoboken (NJ): Wiley; 2011 (https://www.wiley.com/en-us/Medical+Product+Regulatory+Affairs:+Pharmaceuticals,+Diagnostics,+Medical+Devices-p-9783527318773). 29. Wong J, Tong RKY, editors. Handbook of medical device regulatory affairs in Asia. Second edition. Singapore: Jenny Stanford Publishing; 2018 (https://www.jennystanford.com/9789814774314/handbook-of-medical-device-regulatory-affairs-in-asia-second-edition/). 30. Mantus D, Pisano DJ, editors. FDA regulatory affairs. Third edition. Boca Raton (FL): CRC Press; 2014 (https://www.routledge.com/FDA-Regulatory-Affairs-Third-Edition/Mantus-Pisano/p/book/9781841849195).